ΑΠΑΙΤΗΣΕΙΣ ΧΩΡΟΥ ΕΓΚΑΤΑΤΑΣΗΣ ΑΞΟΝΙΚΟΥ ΤΟΜΟΓΡΑΦΟΥ 16 ΤΟΜΩΝ

ΤΥΠΟΥ A  (Aυξημένων Δυνατοτήτων)

ΕΜΒΑΔΟΝ

ΣΤΑΤΙΚΟΤΗΤΑ

The floor construction has to be performed solid and free of vibration, e.g. concrete flooring C20/25 to C50/60 corresponding to DIN EN 206-1, according to the maximum values as specified in the text block “Floor and building vibrations”. ESD flooring is not required for CT systems and CT options. An on-site  subconstruction has to fulfil the flooring conditions. It is recommended to test the weight capacity of the concrete or composite flooring by a stress analyst.

Fastening the gantry on the floor is possible but only necessary in countries prone

to earthquakes or according to local regulations. Securing the patient table to the floor is mandatory. Please refer to local regulations.

The maximum static floor load on the measuring points A to D due to the gantry's own weight is each 5995 N .  For a static interpretation, a maximum value of 5995 N (3 point surface) has to be factored in.

The maximum amplitude for the dynamic floor load

(gantry rotation) is each 250 N.

The bearing surface is each 13 cm².

ΗΛΕΚΤΡΟΛΟΓΙΚΕΣ ΑΠΑΙΤΗΣΕΙΣ

ΑΠΑΙΤΗΣΗΣ ΨΥΞΗΣ