## Sheet1
| Biosensors International |  | Terumo |  | Abbott |  |  |  | Medtronic |  | Boston Scientific |  |  |  | Biotronik | Merillife
| BioMatrix Flex | BioMatrix NeoFlex | Nobori | Ultimaster | Xience V | Xience Pro | Xience Prime | Xience Xpedition | Endeavor | Resolute Integrity | Promus | Promus Element | Promus Premier | Synergy | Orsiro | Biomime
Technical Information
Number of references | 44 | 44 | 25 |  | 36 |  |  | 54 | 36 | 36 | 30 | 42 | 94 | 48 | 54
Stent lengths (mm) | 8 / 11 / 14 / 18 / 24 / 28 mm 33 & 36 mm (except for 2,25 and 4,00 mm diameter) | 8 / 11 / 14 / 18 / 24 / 28 mm 33 & 36 mm (except for 2,25 and 4,00 mm diameter) | 8/14/18/24/28 |  | 8/12/15/18/23/28/33/38/48 |  | 8, 28, 33, 38 | 8 / 12 / 15 / 18 / 23 / 28 mm 33 / 38 (except for 2,25mm diameter) | 8-9/12/14-15/18/24/30 | 8-9/12/14-15/18/24/30/34/38 | 8/12/15/23/28 | 8/12/20/24/28/32/38 | 8/12/16 / 20/24/28/32/38 | 8/12/16/20/24/28/32/38 | 9 / 13 / 15 / 18 / 22 / 26 / 30 / 35 / 40 | 8/13/16/19/24/29/32/37/40/44/48
Stent diameters (mm) | 2,25 / 2,50 / 2,75 / 3,0 / 3,50 / 4,00 | 2,25 / 2,50 / 2,75 / 3,0 / 3,50 / 4,00 | 2,25 / 2,50 / 2,75 / 3,0 / 3,50 |  | 2,25/2,5/2,75/3,0/3,5/4,0 |  | 2.25, 4.0, | 2,25 / 2,5 / 2,75 / 3,00 / 3,25 / 3,5 /4,00 | 2,25/2,5/2,75/3,0/3,5/4,0 | 2,25/2,50/2,75/3,00/3,50/4,00 | 2,25/2,5/2,75/3,0/3,5/4,0 | 2,25/2,50/2,75/3,00/3,50/4,00 | 2,25/2,50/2,75/3,00/3,50/4,00 | 2,25/2,50/2,75/3,00/3,50/4,00 | 2,25 / 2,50 / 2,75 / 3,00 / 3,5 / 4,00 | 2.00/2.25/2.50/2.75/3.00/3.50/4.00/4.50
Shaft design | Hypotube | Hypotube
Shaft coating | PTFE | Polymide Jacket
Shaft length | 142 cm | 142 cm | 145 cm |  |  |  |  | 145 cm |  | 140 cm |  |  |  |  | 140 cm | 140-142 cm
Balloon Compliance | Semi-compliant | Semi-compliant | low compliant
Markers | 2 | 2 | 2 |  |  |  |  | 2 |  | 2 |  | 2 | 2 | 2 | 2 | 2
Lesion Entry Profile | 0.0182" / 0.46 mm | 0.0164” based on bench test results | 0,0169'' / 0,43mm |  |  |  |  | 0.0212" |  | 0,0216" for a 3,00 mm |  | 0,018" | 0,018" | 0,018" (0,46mm) | 0,017" - claimed on the brochure 0,0238 - internal benchtest
Crossing Profile (3.00 mm) | 0,045" / 1,14 mm (for length 8 to 28mm) 0,047" / 1,19mm (for length 33 & 36) * see DES extra info tab | 0,045" / 1,14 mm (for length 8 to 28mm) 0,047" / 1,19mm (for length 33 & 36) * see DES extra info tab | 0,044'' / 1,11 mm for a 3,0mm * see DES extra info tab |  |  |  |  | 0,0425" |  |  |  |  | 0,040" |  | 0,99 mm for 3,0mm | 0.95-0.99 mm / 0.037”-0.039” * see DES extra info tab
Metal to Artery ratio (%) | see DES extra info tab | see DES extra info tab | [12 empty] | see DES extra info tab
Stent Material | Stainless steel 316 L | Stainless steel 316 L | Stainless steel | Cobalt Chromium | Cobalt Chormium |  | Cobalt Chromium | L-605 Cobalt Chromium | Cobalt Chromium | Cobalt Chromium | Cobalt Chromium | Cobalt Chromium |  | Platinum Chromium | Cobalt Chromium | Cobalt Chromium
Stent Design | Corrugated rings | Corrugated rings
Stent Strut thickness | 0,0047"±0,0003’’/ (0,12 mm) 120µm (119 µm in internal benchtests) | 0,0047'' (0,12mm) 120µm (119 µm in internal benchtests) | 0,0047'' / 0,12mm / 120µm | 80 μm | 0,0032'' / 0,081 mm |  |  | 81µm/ 0,0032" | 0,091 mm | 89 µm / 0,0035" 90 µm (in internal benchtests) | 0,081 mm | 81µm (0,081 mm) / 0,0032" |  | 74µm | 60µm for 3.0 0,0024" | 65 µm (0.065mm or 0.0026")
Stent strut length | 1,2 mm | 1,2 mm
Stent strut width | 1,2 for tehe 6 crown model 1,0 for the 9 crown model | 1,2 for tehe 6 crown model 1,0 for the 9 crown model
Cell opening after dilatation | 1,56 mm (for 3,0 mm) | 1,56 mm (for 3,0 mm) |  |  | 1,21mm |  |  |  | 1,19mm |  | 0,21mm
Max Cell opening | 3,0 mm stent = 3,34 mm 3,5 mm stent = 4,19 mm | 3,0 mm stent = 3,34 mm 3,5 mm stent = 4,19 mm | 3,5 mm stent = 3,72 mm | [6 empty] | 2.25-2.75: up to 3.50 mm / 3.00-4.00: up to 4.75 mm |  |  |  |  | 13.8 mm
Elastic recoil | 2.46% | ≤ 5% |  |  | 4.86% |  |  |  | 4.6% | 4.91 | 4.86% | 3.64% |  |  | 4,7 % (for 2,25 to 3mm) 4,6% (for 3,5 & 4,0 mm) | 3%
Foreshortening | 0.39% | ≤ 10% | [7 empty] | 0.48% | [5 empty] | 0.29%
Radial Strength | > 0.67 bar / 500 mmHg | > 0.67 bar / 500 mmHg
Nominal Pressure | 6 atm | 6 atm | 8atm |  |  |  |  | 10 atm |  | 9 atm |  |  |  |  | 9 atm | 9 atm
Rated Burst Pressure | 16 atm (for 2.25 - 3.00 mm) 14 atm (for 3.50 - 4.00 mm) | 16 atm (1621 kPa) 2.25-3.00 mm 14 atm (1418 kPa) 3.50-4.00 mm | 16 atm (1621 kPa) - 2.5 to 3.0 mm 14 atm (1418 kPa) - 3.5 mm |  |  |  |  | 18 atm |  | 16 atm (2.25 -3.5mm) 15 atm (4.0mm) |  |  |  |  | 16 atm - 2.0 to 4.0 mm 14 atm - 4.5 to 5.00 mm | 14/16 atm
Guiding catheter compatibility | 5F (for 2.25 - 3.00 mm) 6F (for 3.50 - 4.00 mm) | 5 F | 5F |  | 5F |  |  | 5F (for stent 2,25 to 3,5 mm) 6F (for stent 4,0 mm) | 5F | 5F | 5F |  |  |  | 5F | 5F
Guidewire compatibility | 0.014"/ 0.36 mm | 0.014"/ 0.36 mm | 0,014'' / 0,36 mm |  |  |  |  | 0,014" |  | 0,014" |  |  |  |  | 0,014"
Drug Name | BA9 - Biolimus A9 | BA9 - Biolimus A9 | BA9 - Biolimus A9 | Sirolimus | Everolimus |  | Everolimus | Everolimus | Zotarolimus | Zotarolimus | Everolimus | Everolimus | Everolimus | Everolimus | Sirolimus | Sirolimus
Drug Dosage | 15,6 μg/mm stent length | 15,6 μg/mm stent length | 15,6 μg/mm stent length | 3,9 μg/mm sten | 100μg/mm2 |  | 100μg/mm2 | 100 μg/cm2 | 10μg/mm2 | 1,6μg/mm2 | 1μg/mm2 |  |  |  | 1,4 μg/mm2 | 1,25μg/mm2
Coating thickness | 11μm | 11μm | 11μm | 15 μm | 7,8 μm |  | 7,8 μm | 7,8 μm | 4,8 μm | 4,8 μm | 7,8 μm | 7,8 μm |  | 4 μm |  | 2μm
Polymer | Biodegradable Polymer | Biodegradable Polymer | Biodegradable Polymer |  | Durable Polymer |  |  | Fluorinated Copolymer | Durable Polymer | Durable Polymer | Durable Polymer | Durable Polymer |  | Biodegradable Polymer | BIOlute
Coating configuration | Abluminal | Abluminal | Abluminal |  |  |  |  | Circumferential | [5 empty] | Abluminal
Shelf life | 24 months | 24 months | [7 empty] | 18 months